## Slide 1
Self-Built Nitride MBE System
Project proposal
30 July 2026

## Slide 2
In-house built machine
NITRIDE MBE PROPOSAL · MOTIVATION
01
Cost
Estimated US$0.9–1.5M for entire chamber; Depends on quotes; annual run cost ~US$60–120k
02
Capability & Needs
1200 °C substrate heater for selective-area regrowth and normal epitaxy
03
In-Situ Control
RHEED video output; Heater temperature knobs; Shutter & flux control
OBJECTIVE
MBE chamber + plasma source + heater + effusion cells + control software — one machine for epitaxy, selective area growth, and autonomous operation

## Slide 3
MBE components
NITRIDE MBE PROPOSAL · PARTS
MBE chamber — Veeco GEN II / Riber 32 class
Growth, buffer and load-lock chambers · chamber frame · transfer system · liquid nitrogen shroud · load-lock cryopump · racks
RF N₂ plasma source
600 W · 13.56 MHz ·RF coil. matching network
N₂ gas delivery system
99.9999% pure supply · getter purifier · Mass flow controller
1200 °C heater
PBN-PG heater
Effusion cells ×6
Ga · cold-lip Al · Cu · Si · Mg · spare
Maglev turbomolecularpump
≥2000 L/s N₂ + dry backing pump
Ion pump + Titanium sublimation pump
Valved: base-pressure duty only
RHEED
gun · phosphor screen · camera
Band-edge pyrometer
true wafer temperature reading
Residual gas analyzer + ion gauges
incl. retractable beam-flux monitor
Power supply + Proportional integral derivative
one DC supply + control loop per heater zone
LN₂ distribution
phase separator · vacuum-insulated lines
Bake-out heater · water-loop chiller · Uninterruptible Power Supply

## Slide 4
Cost Estimate
NITRIDE MBE PROPOSAL · BUDGET
01
GEN II / Riber 32 class, refurbished
used chassis
0.20–0.50
02
Plasma source · 1200 °C manipulator · effusion cells
nitride core
0.21–0.42
03
Maglev turbo · RHEED 30 keV · band-edge pyrometry · RGA
vacuum + diagnostics
0.19–0.32
04
Software built in-house on commercial interlocks
controls + DAQ
0.04–0.08
05
LN₂ · 3-phase power + UPS · water · bake-out
facility
0.15–0.26
06
Leak-chase · calibration · spares
install + contingency
0.10–0.25
US$M
TOTAL
US$0.9–1.5M · run cost ~US$60–120k / yr

## Slide 5
Three risks we manage from day one
NITRIDE MBE PROPOSAL · RISKS
01
Export control
MBE growth equipment is a listed strategic commodity — a Hong Kong import licence (3B001(a)(3)) is required whatever the origin. US-origin parts additionally need a BIS licence, since the US has treated HK as mainland China since Dec 2020, and approval is not assured.
MITIGATION
classify every candidate part's ECCN before any PO; prefer EU sources (Riber, MBE-Komponenten, STAIB) or domestic builders as fallback; licence path confirmed = go / no-go gate P0
02
Temperature spec
Selective-area growth only turns on above ~930–940 °C on-wafer, and AlN wants more. Used chassis come with ~900 °C GaAs manipulators — unusable — and retrofitting later means re-opening the chamber. Heater filaments also degrade in nitrogen plasma.
MITIGATION
order the 1200 °C manipulator as the first long-lead item; specify a W or PBN/PG heater for the N-plasma environment; treat filaments as consumables and stock spares
03
N₂ gas load
Nitride growth runs at ~10⁻⁵ Torr of continuous N₂. Ion pumps saturate and cryopumps need constant regeneration at this load — a mis-sized vacuum stack is usually discovered only at first plasma, after the rebuild is finished.
MITIGATION
maglev turbo ≥2000 L/s as the primary growth pump, sized up front — never retrofitted; valved ion pump + TSP kept for idle / base pressure; getter-purified N₂ protects film purity